## Slide 1: Test Methods, Validation & Verification
ISO/IEC 17025:2005

## Slide 2: Objectives; At the end of this course participants will be able to…
Participant will understand that only the most current, approved and appropriate method should be used.
Participant will comprehend the requirements for validating a non-standard or in-house method.

## Slide 3: Abbreviations & Acronyms
Accuracy-(Method Bias) nearness to the true value; can be determined by one of the following techniques:
  Use of certified reference materials;
  Use of reference method of known uncertainty
  Use of recovery from spiked samples
  Reference material, if available, should be carried through entire procedure with each batch of test samples. If fortified or spiked samples are used, the method of fortification should be described.
Bias-systematic error manifested as a consistent unidirectional deviation from the true value
Calibration-the comparison of two measurement devices or systems; one of known uncertainty (your standard) and one of unknown uncertainty (your test equipment)
Certified Reference Material (CRM)-reference material that has documentation by an authoritative body and providing one or more specified property values with associated uncertainties and traceabilities
Laboratory developed methods-methods that are developed using “in-house” procedures, or by changing an established reference procedure.

## Slide 4: Abbreviations & Acronyms
Reference method-methods that have been published or documented by an authorizing body
SOP-Standard Operating Procedure
Validation-the confirmation by examination and the provision of objective evidence that the particular requirements for a specific intended use are fulfilled
Verification-laboratory verifies by an established process that they are capable of using a validated method for its intended use

## Slide 5: Some distinctions
Policies
“What we do”
Processes “How it happens”
Procedures and Job Aids
“How to do it”
Forms & Records “What happened”

## Slide 6: Methods
The laboratory shall use appropriate methods that meet the needs of the customer and use procedures for all tests that should include these items:
Sampling
Handling
Transport
Storage
Sample Preparation
Estimation of measurement uncertainty
Statistical techniques for analysis

## Slide 7: When Selecting a Method
It should be approved or published in
International, regional or national standards
Reputable technical organizations
Relevant scientific journals
Specified by manufacturer of equipment
Use the latest valid edition unless it is not appropriate or possible to do so
Supplement with additional details when necessary to ensure consistent application
The laboratory must confirm they can properly operate any standard method

## Slide 8: When Selecting a Method
Documented
Technically justified
Authorized and accepted by the customer
Appropriate for the test
Laboratory must confirm that it can properly operate the method before introducing the test. If the method changes, the confirmation must be repeated.
Deviations shall occur only if the deviation has been

## Slide 9: Laboratory Developed Methods
Must be appropriate for intended use
Customer must be informed
Method is validated:
Validation is the confirmation by examination and the provision of objective evidence that the particular requirements for a specific intended use are fulfilled.
It is a planned activity
It is assigned to qualified personnel with adequate resources
Follow the (Insert name here) Validation and Verification SOP (standard operating procedure)
Record the results obtained, the procedure for the validation, and a statement as to whether the method is fit for the intended use
Includes processes for sampling, handling, and transportation
Laboratory shall validate these methods
Non-standard
Laboratory designed/developed
Standard methods outside their intended scope (any amplifications or modifications)

## Slide 10: Validation
Techniques used should include one or more of the following:
Calibration using reference standards or materials
Comparison of results with other methods
Interlaboratory comparisons
Systematic assessment of the factors influencing the result
Assessment of the uncertainty of the results based on scientific understanding of the theoretical principles of the method

## Slide 11: New Test Methods Should Contain:
Appropriate identification
Scope
Description of type of item to be tested
Parameters of quantities and ranges to be determined
Apparatus and equipment, including technical performance requirements
Reference standards and material required
Environmental conditions required and any stabilization period needed

## Slide 12: New Test Methods Should Contain:
Description of procedure
Identification, handling, transporting, storage, and preparation of items
Checks to be made before work is started
Checks that equipment is working properly, required calibration and adjustments are made prior to each use
Method of recording observations and results
Safety measures to be observed
Criteria and requirements for approval/rejection
Data to be recorded and method of analysis and presentation
Uncertainty or the procedure for estimating uncertainty

## Slide 13: Validation
Range and accuracy of the values obtained should include:
Uncertainty of the results
Detection limit
Selectivity of the method
Linearity
Limit of repeatability and/or reproducibility
Robustness against external influences and/or cross-sensitivity against interference from the sample matrix
Assessed for the intended use
Relevant to the customer

## Slide 14: Validation, continued
When changes are made to a validated method, the influence of the changes should be documented, and if appropriate, a new validation carried out
Range and accuracy of the values (uncertainty)
These items shall be relevant to the customer’s needs:
Detection limit
Selectivity of the method
Linearity
Limit of repeatability and/or reproducibility
Robustness against external influences
And/or cross-sensitivity against interferences from the matrix of the sample

## Slide 15: Validating Data/Software
Calculations (manual and electronic) and data transfers shall be subject to appropriate checks in a systematic manner
When computers or automated equipment are used for the acquisition, processing, recording, reporting, storage, or retrieval of data the laboratory shall ensure that:
  Computer software developed by the user is documented in sufficient detail and is suitably validated as being adequate for use
    Excel spreadsheets with calculations only (documents)—Lock the formula cells
    Excel spreadsheets with results after calculations (records)—pdf the record so no changes can be made
  Procedures are established and implemented for protection of data (integrity and confidentiality, storage, transmission, and processing)
    Commercial software (e.g., word processing, database, statistical programming) is considered sufficiently validated
    Laboratory software and modifications to commercial software must be validated before use

## Slide 16: Results Must be Reported
Accurately
Clearly
Unambiguously
In accordance with specific test instructions
With information requested by customer & necessary to interpretation

## Slide 17: Test Reports Shall Contain
Title, e.g., “Test Report”
Name and address of laboratory
Unique identification of the test report and clear numbering system, e.g., page 1 of 3
Name and address of customer
Identification of method used
Description, condition, and identification of item being tested
Date of receipt
Date test was performed
Reference to sampling plan/procedures used, where relevant
Test results with units of measurement
Name, function, and signature of person(s) authorizing the test report
A statement to the effect that the results relate only to the items tested

## Slide 18: Test reports shall contain, where necessary to interpretation of results
Deviations from, additions to, or exclusions from the test method, and information on specific test conditions, such as environmental conditions
A statement of compliance/non-compliance with requirements and/or specifications
A statement on the estimated uncertainty of measurement
  Where relevant to validity or application of test results
  When a customer requires it
  When uncertainty affects compliance to a specification limit

## Slide 19: References
ISO/IEC 17025:2005
  https://www.iso.org/standard/39883.html