## Slide 1: Traceability

## Slide 2: Objectives
At the end of this learning activity, the participant will be able to:
Define the terms calibration, certified reference material and measurement traceability mean in relation to laboratory testing
Identify the difference between reference standards and reference materials
Describe the items to document in the laboratory to meet traceability requirements

## Slide 3: Some Distinctions
Policies
“What we do”
Processes
“How it happens”
Procedures and Job Aids
“How to do it”
Forms & Records “What happened”

## Slide 4: Terms and Acronyms
Certified Reference Material (CRM) - Reference material that has documentation by an authoritative body and providing one or more specified property values with associated uncertainties and traceabilities
CGPM - Conférence Générale des Poids et Mesures (French: General Conference on Weights and Measures)
CIPM - Comité International des Poids et Mesures (International Committee of Weights and Measures)
ISO Guide 34 - General Requirements for the competence of reference material producers

## Slide 5: Terms and Acronyms
NIST - National Institute of Standards and Technology
PT – Proficiency Test
QASP - Quality Assurance Standard Procedure
QC - Quality Control
QMS - Quality Management System
Reference Standard - Hard and not consumable items, such as weights and thermometers, used for calibration measurements that can be traced to a standard unit of the SI (International System of Units)
SI - International System of Units
SOP - Standard Operating Procedure

## Slide 6: What is Traceability?
Traceability is the property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty
Traceability to the SI (International System of Units) -metrological traceability to a measurement unit of the International System of Units
Only measurement results can be traceable

## Slide 7: What do I Need to Support a Claim of Traceability?
*Document the measurement process or system used and provide a chain of calibrations used to establish connection to the reference. Common elements include:
  a clearly defined particular quantity that has been measured (e.g. weight, volume, temperature, etc.)
  a complete description of the measurement system or working standard used to perform the measurement (what is the measurement based on or compared to: did you use a CRM, a NIST traceable thermometer; a calibrated pipette?)
  a stated measurement result, which includes a documented uncertainty (when items have been calibrated or manufactured to certain accredited bodies, they should include an uncertainty measurement)
  *From NIST website FAQ I.A.5:
  https://www.nist.gov/traceability/supplementary-materials-nist-policy-review
Notes: What do I need to do to support a claim of traceability?
To support a claim, the provider of a measurement result must document the measurement process or system used to establish the claim and provide a description of the chain of calibrations that were used to establish a connection to a particular specified reference. There are several common elements to all valid statements or claims of traceability:
a clearly defined particular quantity that has been measured a complete description of the measurement system or working standard used to perform the measurement a stated measurement result, which includes a documented uncertainty a complete specification of the reference at the time the measurement system or working standard was compared to it an 'internal measurement assurance' program for establishing the status of the measurement system or working standard at all times pertinent to the claim of traceability an 'internal measurement assurance' program for establishing the status of the specified reference at the time that the measurement system or working standard was compared to it
An internal measurement assurance program may be quite simple or very complex, the level or rigor to be determined depending on the level of uncertainty at issue and what is needed to demonstrate its credibility. Users of a measurement result are responsible for determining what is adequate to meet their needs.
For information and guidance on expressing measurement uncertainty, see http://physics.nist.gov/cuu/.
See Section V Checklist for Traceability through Calibration for additional guidance.

## Slide 8: Metrology Systems
ISO/IEC 17025:2005
Requires laboratories to have calibration systems in place that ensure that their measurements (results) are comparable with any other laboratory.
ISO Guide 34
The National Institute of Standards and Technology (NIST) was established as the National Metrology Institute (NMI) for the US by Congress in 1901.
“Specifies general requirements in accordance with which a reference material producer has to demonstrate that it operates, if it is to be recognized as competent to carry out the production of reference materials.” http://www.iso.org/iso/catalogue_detail.htm?csnumber=50174
NIST

## Slide 9: NIST and traceability
Requires the establishment of an unbroken chain of calibrations to specified references.
Test equipment calibration is not traceable to the NIST, but to the SI through NIST: http://www.nist.gov/index.html
Traceability
Metrological traceability
NIST
Responsible for developing, maintaining and disseminating national standards - realizations of the SI - for the basic measurement quantities, and for many derived measurement quantities.
NIST
Responsible for assessing the measurement uncertainties associated with the values assigned to these measurement standards.

## Slide 10: What does it Mean to be “Traceable to NIST’?
Reference measurement results can be traced through an unbroken chain of calibrations, including determining the uncertainties at each step, to NIST standards as the specified references.
These references may be, for example, standards developed and maintained by NIST:
  broadcast signals controlled or monitored by NIST (such as standard time and frequency signals)-used to measure timers, for example,
  NIST Standard Reference Materials (SRMs), or
  NIST-Traceable Reference Materials (NTRMs).
    The chain of calibrations may be short, if the user has instruments or artifacts calibrated by NIST or acquires standards from NIST and references measurement results to those.
    It may be longer, if the user references other calibrations in a chain of calibrations back to stated references developed and maintained by NIST.
Notes: To achieve traceability of measurement results to standards maintained by NIST, you need to reference your measurement results through an unbroken chain of calibrations, including determining the uncertainties at each step, to NIST standards as the specified references. These references may be, for example, standards developed and maintained by NIST: broadcast signals controlled or monitored by NIST (such as standard time and frequency signals), NIST Standard Reference Materials (SRMs), or NIST-Traceable Reference Materials (NTRMs). The chain of calibrations may be short, if the user has instruments or artifacts calibrated by NIST or acquires standards from NIST and references measurement results to those. It may be longer, if the user references other calibrations in a chain of calibrations back to stated references developed and maintained by NIST. Also see Question I.A.5 above and Section V Checklist for Traceability through Calibration for additional guidance.

## Slide 11: Reference Standards vs. Materials
Hard and non-consumable items
  NIST traceable thermometers
  Weights
These items should be used for calibration only to maintain accuracy and integrity
They should be calibrated by an ISO/IEC 17025 accredited body that can provide traceability
Consumables
  Liquid standards
  Proficiency testing materials
If they cannot be traced to SI units, then they should be provided by a competent supplier to give a reliable physical or chemical characteristic of a material (e.g., PTs, CRMs)
Reference Standards
Reference Materials

## Slide 12: What is Calibration?
Calibration is the comparison of two measurement devices or systems, one of known uncertainty (reference standard) and one of unknown uncertainty (test equipment)

## Slide 13: What is Calibration?
A response to a known value which establishes a relationship between the known value and an indicator and established tolerance
  I have the result, I know what it should be
  What is my answer in relation to the known value and how much can difference I (and my customer) can tolerate?
Tolerance (+ or -)
Known Value
End result
Measured value with some Uncertainty
Measured again with some Uncertainty

## Slide 14: Calibration Laboratories
Establish traceability of its own measurement standards and measuring instruments to the International System of Units (SI) (e.g., NIST, CIPM, CGPM)
Issue calibration certificates that shall contain the measurement results, including the measurement uncertainty
Use vendors that have been accredited to ISO Guide 34

## Slide 15: Traceability Includes Specific Requirements for
Calibration & QC Materials
Testing
Reagents, Media, Cultures
Reference Standards
Reference Materials
Intermediate Checks
Transport
Critical & Support Equipment
Chemicals

## Slide 16: What is Traceability? A Calibration Hierarchy For example: Reference Standards shall be calibrated by a body that can provide traceability (traced back) to a National Standard of Measurement (a measurement unit, procedure, or standard)
Stock Solution
Result
Secondary reference standard
Working solution
Reference to International or National Measurement Standard, e.g., NIST, CIPM
Primary reference standard

## Slide 17: Intermediate Checks
There need to be defined procedures and schedules for carrying out checks needed to maintain confidence in the calibration status of:
Reference standards
Primary standards
Working standards
Transfer standards
Reference materials

## Slide 18: Testing Laboratories
Traceability applies to measuring and test equipment unless you can prove that it is a minor component of measuring uncertainty. In this case, you must provide the uncertainty of the instrument.

## Slide 19: Testing Laboratories
Where traceability to SI units is not possible or relevant, the laboratory should
Use Certified Reference Materials provided by a competent supplier to give a reliable physical or chemical characterization of a material
Use specified methods and/or consensus standards that are clearly described and agreed by all parties concerned

## Slide 20: Methods
5.4.2: The laboratory shall use appropriate methods that meet the needs of the customer and are appropriate for tests and use procedures for all tests that should include these items:
Sampling
Handling
Transport
Storage
Sample Preparation
Estimation of measurement uncertainty
Statistical techniques for analysis

## Slide 21
Procedures written and Information you need to keep on:
Reagents, Reference Materials and Reference Standards
Manufacturer
Purity
Date of Receipt
Storage conditions
Expiration
Prepared by
Dates of Use
Certificate of Authenticity
Transport
How to Purchase
Unique Identifier

## Slide 22: Key Questions to Ask in the Laboratory
Has the laboratory identified all the measuring equipment that would affect the validity of measurements?
Has it been properly calibrated in a manner that provides traceability to the international measurement system?
Does the QMS address that calibration is maintained and monitored? Is drift detected?
Does the laboratory keep records of the calibration status of the equipment? Would an auditor be able to trace the calibration status to any point in the past?
Are all reagents traceable to a CRM or other acceptable source (such as a consensus standard)?

## Slide 23: Key Questions to Ask in the Laboratory
Does the laboratory hold a master reference thermometer?
  If so, did the thermometer come with a calibration certificate issued by an accredited calibration laboratory? Is it in need f recalibration?
Does the laboratory have a written procedure and training program for carrying out calibration of working thermometers with adequate applied quality controls?
Are there records that show an auditor that these criteria are met on a routine basis?

## Slide 24: Summary
Traceability is an unbroken record of documentation of measurement that can be related to stated references
  Tracing reagents back to their source (stock or CRM)
  Equipment used (pipettes, thermometers, balances)
  Including their associated uncertainties

## Slide 25: References
ISO/IEC 17025-www.iso.org/39883.html
NIST-http://www.nist.gov/
NIST Mercury Thermometer Alternatives: What Is Traceability?- http://www.nist.gov/pml/mercury_traceability.cfm
ISO Guide 34- http://www.iso.org/iso/catalogue_detail.htm?csnumber=50174 ‘
Complying with ISO 17025: A Practical Guidebook for meeting the requirements of laboratory accreditation schemes based on ISO 17025:2005 or equivalent national standards by the United Nations Development Organization, Vienna, 2009 https://www.unido.org/fileadmin/user_media/Publications/Pub_free/Complying_with_ISO_17025_A_practical_guidebook.pdf