## Slide 1: Traceability
QMS Quick Learning Activity
www.aphl.org

## Slide 2: Objectives
At the end of this learning activity, the participant will be able to:
Define the terms calibration, certified reference material, and traceability in relation to laboratory testing
Identify the difference between reference standards and reference materials
Describe the items to document in the laboratory to meet traceability requirements

## Slide 3: Terms and Acronyms
Certified Reference Material (CRM) - Reference material that has documentation by an authoritative body and providing one or more specified property values with associated uncertainties and traceabilities
CGPM - Conférence Générale des Poids et Mesures (French: General Conference on Weights and Measures)
CIPM - Comité International des Poids et Mesures (International Committee of Weights and Measures)
ISO 17034 - General Requirements for the Competence of Reference Material Producers

## Slide 4: Terms and Acronyms
NIST - National Institute of Standards and Technology
PT – Proficiency Test
QC - Quality Control
QMS - Quality Management System
Reference Standard - Substance or item prepared for use as a standard in an assay, identification, or purity test.
SI - International System of Units

## Slide 5: What is Traceability?
Traceability is the property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty.
Traceability to the SI means metrological traceability to a measurement unit of the SI.
Only measurement results can be traceable.

## Slide 6: What do I Need to Support a Claim of Traceability?
*Document the measurement process or system used and provide a chain of calibrations used to establish connection to the reference.
Common elements include:
a clearly defined quantity that has been measured (e.g. weight, volume, temperature)
a complete description of the measurement system or working standard used to perform the measurement (What is the measurement based on or compared to: Did you use a CRM, a NIST-traceable thermometer; a calibrated pipette?)
a stated measurement result, which includes a documented uncertainty. When items have been calibrated or manufactured to certain accredited bodies, they should include an uncertainty measurement.
*From NIST website FAQ I.A.5: https://www.nist.gov/traceability/supplementary-materials-nist-policy-review

## Slide 7: Metrology Systems
ISO/IEC 17025:2017
Requires laboratories to have calibration systems in place that ensure that their measurements (results) are comparable with any other laboratory.
ISO 17034
Specifies general requirements for the competence and consistent operation of reference material producers.
Sets out the requirements in accordance with which reference materials are produced. It is intended to be used as part of the general quality assurance procedures of the reference material producer.
Covers the production of all reference materials, including certified reference materials. https://www.iso.org/standard/29357.html
NIST
The NIST provides technology, measurement, and standards for the United States.

## Slide 8: NIST and traceability
Traceability
Test equipment calibration is not traceable to the NIST, but rather to the SI through NIST
Metrological traceability
Requires the establishment of an unbroken chain of calibrations to specified references.
NIST
Responsible for developing, maintaining and disseminating national standards - realizations of the SI - for the basic measurement quantities, and for many derived measurement quantities.
NIST
Responsible for assessing the measurement uncertainties associated with the values assigned to these measurement standards.

## Slide 9: What does it Mean to be “Traceable to SI”?
Reference measurement results can be traced through an unbroken chain of calibrations, including determining the uncertainties at each step, to SI units.
There are seven base units in the SI system:
  Mass - kilogram (kg)
  Time - second (s)
  Temperature - kelvin (K)
  Electric current - ampere (A)
  The amount of a substance, mole (mol)
  Luminous intensity - candela (cd)
  Length - meter (m)
The SI is explained more in the following brochure (2019): https://www.bipm.org/utils/common/pdf/si-brochure/SI-Brochure-9.pdf

## Slide 10: Use of Reference Material to be Traceable to the SI
Reference materials are materials or substances, one or more of whose property values are sufficiently homogeneous and well established to be fit for its intended use in measurement or in examination of nominal properties.
Use CRMs when possible
    Reference material that has documentation by an authoritative body and providing one or more specified property values with associated uncertainties and traceabilities.
  NIST produces reference materials traceable to SI units that can be purchased for use in laboratories to maintain traceability.
    NIST Standard Reference Materials (SRMs)
    NIST-Traceable Reference Materials (NTRMs)

## Slide 11: Reference Standards vs. Materials
Non-consumable items
  NIST traceable thermometers
  Weights
These items should be used for calibration only to maintain accuracy and integrity
They should be calibrated by an ISO/IEC 17025 accredited entity that can provide traceability
Consumables
  Liquid standards
  PT materials
If they cannot be traced to SI units (CRM), then they should be provided by a competent supplier to give a reliable physical or chemical characteristic of a material (e.g., PT material, other provider)
Reference Standards • Reference Materials

## Slide 12
What is Calibration?
Calibration is the comparison of a measurement device or system under test with one of known uncertainty (limits which contain the true value at specified confidence limits).

## Slide 13: What is Calibration?
Calibration is a process in which a physical measurement or physical response is compared to a known standard
What is my answer, plus or minus my error, in relation to the known value, plus or minus established tolerance?
How much error can by customer tolerate?
Tolerance (+ or -)
End result

## Slide 14: Calibration Laboratories
Establish traceability of its own measurement standards and measuring instruments to SI units (e.g., NIST, CIPM, CGPM)
Issue calibration certificates that contain the measurement results, including the measurement uncertainty
Use vendors that have been accredited to ISO 17034-check certificate!

## Slide 15: Traceability Includes Specific Requirements for
Calibration & QC Materials
Testing
Reagents, Media, Cultures Critical & Support Equipment Reference Standards Reference Materials Intermediate Checks Chemicals
Transport

## Slide 16: What is Traceability? A Calibration Hierarchy For example: Reference Standards shall be calibrated by a body that can provide traceability (traced back) to a National Standard of Measurement (a measurement unit, procedure, or standard)
Result
Working solution
Stock Solution
Secondary reference standard
Primary reference standard
Reference to International or National Measurement Standard, e.g., NIST, CIPM

## Slide 17: Reexamine traceability periodically
There should be defined procedures and schedules for carrying out regular checks needed to maintain confidence in the calibration status of:

## Slide 18: Testing Laboratories
Traceability applies to measuring and test equipment unless you can prove that it is a minor component of measuring uncertainty. In this case, you must provide the uncertainty of the instrument.

## Slide 19: Testing Laboratories
Where traceability to SI units is not possible or relevant, the laboratory should
Use reference materials provided by a competent supplier to give a reliable physical or chemical characterization of a material
Use specified methods and/or consensus standards that are clearly described and agreed by all parties concerned

## Slide 20: Methods
ISO17025:2017 7.2.1: The laboratory shall use appropriate methods that meet the needs of the customer and are appropriate for tests and use procedures for all tests that should include these items:
Sampling
Handling
Transport Storage
Laboratory Sampling
Estimation of measurement uncertainty
Statistical techniques for analysis

## Slide 21
Procedures written and Information you need to keep on:
Reagents, Reference Materials and Reference Standards
QMS Quick Learnin
g Activity
21
Manufacturer
Certificate of Authenticity
Purity
Date of Receipt
Unique Identifier
Storage conditions
Expiration
Prepared by
Dates of Use
How to Purchase
Transport

## Slide 22: Key Questions to Ask in the Laboratory
Has the laboratory identified all the measuring equipment that would affect the validity of measurements?
Has the equipment been properly calibrated in a manner that provides traceability to the SI?
Does the QMS address that calibration is maintained and monitored? Is drift detected?
Does the laboratory keep records of the calibration status of the equipment? Would an auditor be able to trace the calibration status to any point in the past?
Are all reagents traceable to a CRM or reference material from other acceptable source (such as a consensus standard)?

## Slide 23: Key Questions to Ask in the Laboratory
Is the laboratory able to trace temperature calibrations to the SI unit?
– If used, did the reference thermometer come with a calibration certificate issued by an accredited calibration laboratory? Is it being used within the expiration date of the certificate?
Does the laboratory have a written procedure and training program for carrying out calibration of working thermometers with adequate applied quality controls?
Are there records that show an auditor that these criteria are met on a routine basis?

## Slide 24: Summary
Traceability is an unbroken record of documentation of measurement that can be related to stated references
Tracing reagents back to their source (CRM)
Equipment used (pipettes, thermometers, balances)
  Including their associated uncertainties

## Slide 25: References
ISO/IEC 17025:2017 - https://www.iso.org/standard/66912.html
NIST-http://www.nist.gov/
NIST Mercury Thermometer Alternatives: What Is Traceability?- http://www.nist.gov/pml/mercury_traceability.cfm
ISO 17034- https://www.iso.org/standard/29357.html
Supplementary Materials related to NIST Policy on Metrological Traceability http://members.aoac.org/aoac_prod_imis/Programs/RMG_files/Page322.htm
BIPM: The International System of Units https://www.bipm.org/utils/common/pdf/si-brochure/SI-Brochure-9.pdf