|  | Software Update Evaluation SOP |  |  |  |
| --- | --- | --- | --- | --- |
|  | Document #: | Revision #: | Effective Date: | Page 1 |

Purpose

This procedure provides instructions for evaluating instrument software updates to determine necessary actions to ensure the equipment functions according to established criteria to produce the quality of products and services required by [input user-specific institution].

Scope

This document applies to instruments within the [input department section] section used in laboratory testing that generate data to render a test result.

Related Documents

| Title | Document Control Number |
| --- | --- |
| Software Update Form |  |
| Software Release Notes | include applicable version |

Responsibilities

| Position | Responsibility |
| --- | --- |
| All laboratory staff | Ensure software is updated according to established laboratory guidelines Follow documented software update process |
| Unit Supervisor / Team Lead /SME | Ensure documented procedures for the evaluation of software updates are established Ensure documented procedures are followed Review and approve the Software Update Form |
| Quality Specialist | Ensure documented software update procedures are available to the end user Review and approve the Software Update Form |

Procedure

Any updates to the equipment software require evaluation and approval prior to installation.

Obtain the Software Update Form and the Vendor Software Release Notes.

Complete the Software Update Form

Record the current and new software versions.

Note: There are multiple components to the software; record the current and new version numbers for each.

Document the Vendor workflow(s) currently used in the laboratory.

Review each bullet point in the software release notes and evaluate the following:

Determine if the update affects the workflow used in the laboratory.

Determine if the update potentially affects the data output results.

Evaluate the action required to ensure the equipment functions as expected following the software updates:

| Updates affect the workflow? | Updates potentially affect the data? | Required Action |
| --- | --- | --- |
| No | No | None |
| Yes | No | None |
| Yes | Yes | Verification |

Install the software updates as directed in the release notes.

If verification is required, proceed to step 5.3.g; otherwise proceed to step 5.3.h.

Complete a verification run as described below prior to releasing the equipment back into service.

Using a standard, well-characterized sample previously ran in the laboratory, determine what needs to be run.

If the data obtained with the new software versions are comparable to the data obtained with the prior software versions, no further action is needed.

If the data obtained with the new software versions are not comparable to the data obtained with the prior software versions, conduct a larger study with specific samples to evaluate the performance characteristics of the assay.

Attach additional information as needed (e.g. Release Notes documentation, Verification / Validation data) to the Software Update Form.

Sign, date, and obtain applicable reviews and approvals.

Revision History

| Rev # | DCR # | Change Summary | Date |
| --- | --- | --- | --- |

Approval

Reviewed by: _______________________________ Date: _________________

Software Update Form

(new software, version updates, database additions)

Software to test:

Current software name/version:

New software name/version:

Reason for testing:

Provide rationale including whether the software is new or a replacement system, updated software version, database addition, etc.

Testing method:

Detail the methods used for testing including boundary testing, input testing, and stress testing, as applicable.  Describe the vendor workflow currently being used and how the software update will be affected.

Results:

Explain findings of the comparison study including pass/fail and any additional steps used to resolve discrepancies. Include raw and calculated data, figures, and any other applicable information.

X Set Up By: 		Date Tested:

Y Set Up By:	           Date Tested:

Z Set Up By:	          		Date Tested:

Test result:	(circle)		PASS  /  FAIL

Results Reviewed by: _______________________              Date: ___________

Approval by:  _________________________________              Date: ___________