# Clinical Laboratory Risk Assessment Template:

Ebola Virus Disease Testing

# Introduction

This template is designed to assist laboratories in the development of their risk assessment for Ebola Virus Disease (EVD). It may not be an all-encompassing plan as each facility will have their laboratory specific risk assessment procedures.

Standard precautions have been highly effective in preventing transmission of bloodborne infection in the course of handling blood and other potentially infectious material in the clinical laboratory. Standard precautions should be effective in preventing the transmission of Ebola virus and other viral hemorrhagic fever agents in the clinical laboratory. However, Ebola virus is a high-consequence pathogen, and there has been limited experience handling specimens potentially contaminated with such a high-consequence pathogen in a clinical laboratory using current specimen handling procedures and automated instrumentation. Therefore, this risk assessment is provided for enhanced precautions and personal protective equipment (PPE) in handling specimens from patients who may be at risk of having Ebola virus infection.

Hospitals or clinical laboratories in the United States concerned about a patient with potential EVD exposure should contact their local and/or state health departments and follow their jurisdiction’s procedures for consultation on patient assessments and EVD testing.

APHL recommends that facilities refer to the US Centers for Disease Control and Prevention (CDC) EVD Information for Clinicians in US Healthcare Settings for additional information.

# Assessment Information

| Laboratory Unit/Section |  |
| --- | --- |
| Date of Assessment |  |
| Name of Assessor |  |
| Name of Organism/Agent | Ebola Virus |

# Ebola Virus Disease Testing Overview

| # | Procedure | Potential Hazard(s) | Control/Protection | Additional Information |
| --- | --- | --- | --- | --- |
| A | Package receipt and transfer to testing area | Leaking package | Transfer and open inside a Biological Safety Cabinet (BSC) PPE: nitrile/latex gloves, lab coat, safety glasses | Immediately contact safety officer or other appropriate staff as designated by your facility Disinfect exterior of sealed plastic bag prior to transfer to testing area |
|  |  | Unexpected delivery | Immediately transfer to BSC and contact clinical lab staff/supervisor of unexpected delivery If transporting specimen to another facility, use Category A packaging | Notify key staff of expected package delivery All Category A packages should be opened in certified class II BSC using safety blades |
| B | Transport of specimens between testing areas | Breakage of the specimen container | Inside the facility: Specimens should be transported in a clearly labeled, durable, leak-proof transport container, lined with absorbent material Wipe the outside of the container with an approved disinfectant before leaving the patient bedside Bring container directly to the specimen handling area of the laboratory Outside of the facility: Use a Category A shipping container and IATA and DOT requirements | Decontaminate all surfaces of transport container prior to reuse If transporting in biohazard bags, dispose of in biohazard waste: do not use pneumatic tube transport system |
| C | Preparation of specimens for testing | Aerosolization, splash or splatter | Minimize the number of workers handling the specimens Use BSL-3 PPE: Liquid-impermeable closed front gown OR full body jumpsuit double gloves N95 respirator and goggles or full-face shield (ensure eyes and mucous membranes covered) Don PPE with a trained observer Work should be performed inside a certified class II BSC, following CDC’s acceptable BSC practices: Keep sash at the appropriate level Adjust chair to appropriate height for user Minimize unnecessary movements while working in the BSC Limit the traffic around the BSC Work over a moistened, plastic-backed absorbent pad with appropriate disinfectant Only use pipette tips with barrier filters Have a dedicated rigid waste container and sharps container in the BSC | Prepare disinfectant according to manufacturer’s instructions (e.g., EPA List L: Disinfectants for Use Against Ebola Virus) and wipe all materials down with the disinfect before bringing into BSC No exposed skin inside the BSC Bring all necessary materials into the BSC before starting to work. Verify BSC is operating at acceptable flow rate Minimize use of sharps; dispose of all pipette tips and sharps in the dedicated containers in the BSC Specimens, equipment and all materials must be decontaminated before removing from BSC DO NOT set up any viral cultures Put still-sealed transfer pipets in the BSC, as contamination with bleach or nucleic acids equates to false negatives or positives with the BioFire FilmArray |
| D | Trizol inactivation | Accidental exposure | Perform in BSC Use BSL-3 practices that include the following PPE: Fluid resistant back-closing gown Double gloves N95 respirator Goggles or full-face shield (eyes and mucous membranes covered) Vortex inside BSC; ensure microcentrifuge tube is tightly sealed | Prevent contact with skin, eyes and clothing Wash exposed skin with soap and water immediately. Remove all contaminated clothing or shoes. After Trizol inactivation, specimen is no longer infectious and may be handled using BSL-2 practices. |
|  |  | Trizol reactivity | Do not mix Trizol or Trizol waste with bleach or acids Do not use bleach or acids in BSC while Trizol in use | If accidental mixture of Trizol and bleach/acid occurs: Remove PPE Exit BSL-3 as soon as possible and notify safety officer and supervisor Post “Do Not Enter” sign on BSL-3 suite and do not enter for at least one (1) hour (this time frame will be based on your institutional specific risk assessment) |
| E | Vortexing and Centrifuging | Aerosolizing | Vortex inside BSC; esure microcentrifuge tube is tightly sealed Use sealed head rotor inside the BSC Place specimens in centrifuge cups under BSC and cap. Bring bucket to centrifuge to spin, then return to BSC to open Wipe down tube with appropriate disinfectant | Specimens, equipment and all materials must be decontaminated with approved disinfectant before removing from BSC |
| F | After specimen inactivation and before removal of submitted specimen from BSC | Accidental transfer of contaminated material from the BSC | Wipe all tubes with disinfectant before removing from BSC Place remaining specimen in Ziploc plastic bag; disinfect exterior of bag before removing from BSC. Change gloves | A dedicated waste bag for gloves and other waste is recommended. |
| G | Post extraction BSC decontamination | Contamination of BSC surfaces | Wipe the inside of the BSC with approved disinfectant Remove all PPE and discard into biohazard waste stream | Use 10% bleach disinfectant followed by appropriate contact time Wipe down all surfaces in the BSC with 70% alcohol and allow to air dry. |
| H | Waste | External contamination of waste containers | Disinfect outside of waste containers with approved disinfectant before removal from BSC and BSL-3 | Use 10% bleach disinfectant followed by appropriate contact time Autoclave all PPE used in specimen handling and testing If autoclaving is not possible, review the waste handling protocols with facility management to determine what is required |

# Biological Safety

There are examples included in italics of what information might go in the response column, but this should be tailored to the laboratory’s procedures and inventory.

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| 1 | Biosafety level (BSL) established in this unit | N/A, no BSL |  |
|  |  | BSL-3+ | List BSL-3+ processes e.g., Sample processing including inactivation and extraction steps; Sample storage |
|  |  | BSL-2 | List BSL-2 processes: e.g., PCR (after inactivation) and associated reagent preparation |
|  |  | BSL-1 | List BSL-1 processes: |
| 2 | List parts of testing process that pose a risk of aerosol generation | Vortexing |  |
|  |  | Specimen loading | e.g., Loading into BioFire Cartridge by syringe |
|  |  | Pipetting |  |
|  |  | Centrifuging |  |
|  |  | Opening and closing collection and microcentrifuge tubes |  |
| 3 | List all equipment involved in the use and storage of infectious materials Note: All must have biosafety labels affixed | BSC | has biosafety label |
|  |  | Centrifuge | has biosafety label |
|  |  | Freezers/refrigerators | has biosafety label e.g., Have 2 of each - 1 for specimen, 1 for reagents |
|  |  | Pipettes | has biosafety label e.g., Set is dedicated to Ebola testing |
|  |  | Heat block | has biosafety label |
| 4 | Are buckets with safety caps/cups or aerosol tight rotor lids used when centrifuging infectious materials? | No Yes |  |
| 5 | Is health monitoring performed in this Unit? | No Yes | If yes, please indicate frequency and the process: |
| 6 | Are vaccines recommended for work in this Unit? | No Yes | If yes, please indicate how employees are informed of the vaccines, what vaccines are recommended, etc.: There is no commercially licensed vaccine available for Ebola virus. |
| 7 | List all sharps used and whether they include any safety features | Safety blades | includes safety device feature |
|  |  | Syringe | includes safety device feature |
|  |  | Pipette tips | includes safety device feature |
|  |  | Reagent needles | includes safety device feature |
|  |  | No sharps used |  |
| 8 | Is a Biological Safety Cabinet (BSC) in use? | No Yes | Please indicate if the BSC has been certified within the past year, whether the air vents are not blocked and the sash is in place and operable: |

## Comments

Suggest using a buddy system in BSL-3 Ebola Testing Processing and Extraction.

Laboratory instruments should be maintained as they normally would with any other bloodborne pathogen. A risk assessment should be done before testing specimens from a suspect case on specific automated instruments. Laboratories should consult the manufacturers’ instructions, or the manufacturer directly, on whether there is a need to, and how to, decontaminate the interior surfaces or areas of laboratory instruments.

# Chemical Safety

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| ## Proper Labeling |  |  |  |
| 1 | All containers are labeled with chemical name | Yes No N/A |  |
| 2 | Fire Department Permit is posted on laboratory door | Yes No N/A |  |
| 3 | Chemical inventory is updated | Yes No N/A |  |
| 4 | Materials safety data sheets are accessible to staff | Yes No N/A |  |
| ## Chemical Storage |  |  |  |
| 5 | Incompatible chemicals are segregated | Yes No N/A |  |
| 6 | Flammable liquids are… |  |  |
|  | Stored in rated chemical cabinets | Yes No N/A |  |
|  | Stored in flammable-rated refrigerators/freezers | Yes No N/A |  |
| 7 | Flammable liquids are stored in flammable-rated refrigerators/freezers | Yes No N/A |  |
| 8 | Excessive chemicals are stored in a chemical storage room | Yes No N/A |  |
| 9 | Compressed gas cylinders are stored in the laboratory | Yes No N/A |  |
| 10 | Chemicals are stored at eye-level | Yes No N/A |  |
| 11 | Acids and bases are… |  |  |
|  | Stored in a cabinet | Yes No N/A |  |
|  | Stored in a labeled area | Yes No N/A |  |
|  | Stored free from metals | Yes No N/A |  |
| ## Chemical Fume Hoods |  |  |  |
| 12 | Hoods have been certified within the past year | Yes No N/A |  |
|  | Sash is closed when not in use | Yes No N/A |  |
|  | Exhaust air is not blocked by large equipment or containers | Yes No N/A |  |
|  | Hoods are used for hazardous, toxic or flammable procedures | Yes No N/A |  |

## Additional Comments

# Personal Protective Equipment (PPE)

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| ## Education and Training |  |  |  |
| 1 | Laboratory staff are aware of PPE requirements for this laboratory | Yes No |  |
| 2 | Staff receive an annual PPE competency assessment | Yes No |  |
| ## PPE Care |  |  |  |
| 3 | PPE are appropriately stored in laboratory | Yes No |  |
| 4 | PPE are inspected prior to use and are in good condition | Yes No |  |
| 5 | PPE are not worn in the laboratory area | Yes No |  |
| ## PPE in Use |  |  |  |
| 6 | Safety glasses/goggles | Yes No |  |
| 7 | Facial shields/splash guards | Yes No |  |
| 8 | Disposable laboratory coats | Yes No |  |
| 9 | Nitrile gloves | Yes No |  |
| 10 | Respiratory protection | Yes No |  |
|  | If yes, users are enrolled in a respiratory protection program | Yes No |  |
| 13 | Cryo or autoclave gloves | Yes No |  |
| 14 | Over sleeves/booties/bonnet | Yes No |  |
| 15 | Closed-toe shoes that cover entire foot are worn in laboratory | Yes No |  |

## Additional Comments

# Emergency Preparedness

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| 1 | Emergency contact information is posted | Yes No |  |
| 2 | First aid kit is maintained | Yes No |  |
| 3 | Biological spill kit is maintained | Yes No |  |
| 4 | Staff are aware of occupational injury procedures | Yes No |  |

## Additional Comments

# Documentation and Training

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| 1 | Employee(s) completed right-to-know training | Yes No |  |
| 2 | Employee(s) completed unit-specific training | Yes No |  |
| 3 | Employee(s) completed bloodborne pathogens training | Yes No |  |
| 4 | Employee(s) read and understand safety and health plans | Yes No |  |
| 5 | Door sign is up-to-date and posted | Yes No |  |
| 6 | Laboratory microwaves and refrigerators are labeled with “Not for Food or Drink – Biohazard” | Yes No |  |

## Additional Comments

# Waste Management

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| ## Chemical Waste |  |  |  |
| 1 | Containers are labeled with chemical name and percent of each chemical | Yes No |  |
| 2 | Containers are properly sealed | Yes No |  |
| 3 | Containers are in good condition for transport | Yes No |  |
| ## Biohazard Waste |  |  |  |
| 4 | Broken glass is placed in the appropriate receptacle | Yes No |  |

## Additional Comments

# Engineering Controls

| # | Item | Response | Comments |
| --- | --- | --- | --- |
| 1 | Laminar flow hoods | Yes No |  |
| 2 | Transport containers | Yes No |  |
| 3 | Sharps container | Yes No |  |

## Additional Comments

# At Risk Employees

| Name | Signature | Date |
| --- | --- | --- |

# Assessment Approval

This risk assessment should be reviewed annually or after any major changes (e.g., new facility, new employees, new technology, new method, changes in information for organism/agent, etc.). Minor changes should be recorded under Amendments. Major changes require a new risk assessment to be performed.

Reviews have been carried out on the following dates:

| Title | Printed Name | Signature | Date |
| --- | --- | --- | --- |
| Principal Investigator |  |  |  |
| RO/ARO |  |  |  |
| Section Manager |  |  |  |
| Unit Manager |  |  |  |

Reviewed by:

| Testing Personnel | Printed Name | Signature | Date |
| --- | --- | --- | --- |

# LABORATORY SPECIMEN(S) HANDLING LOG

Use one form for each patient—multiple specimens per patient may use single log form.

Receiver of Specimen(s):

Date Received:  							Time:  					Intake #:

Delivery Method:     FedEx/UPS	 Storeroom	 Department of Health Driver	 Messenger/Name:

Laboratory Number Call for Pick-Up:  							Time Picked Up by Laboratory:

# To be completed by the laboratory:

Picked up from Receiver By:

Purpose:    Patient Clinical Test	 QC		 Training/Validation 	 		  Proficiency

Number of Specimens in Package: 			 	  	Sample Type(s):

Accession #: (1) 		 				 	(2) If Applicable:

Patient ID # on Submission Form: 		 		 	Submitter:

| Function | Responsible Technologist | Signature or Initials of Responsible Technologist | Date and Time |
| --- | --- | --- | --- |
| Delivered to Lab/BSL-3 |  |  |  |
| Accessioned By |  |  |  |
| BSL-3 Processing |  |  |  |
| RNA Extraction |  |  |  |
| PCR Reaction Setup/Master Mix/Template |  |  |  |
| ABI Run |  |  |  |
| ABI Analysis |  |  |  |
| Storage of Original Sample(s) Unit: Location: |  |  |  |
| Storage of RNA Unit: Location: |  |  |  |
| LIMS Result Verification and Report Verification of Report Received |  |  |  |
| Results Reported To: CDC ______ Specimen(s) Shipped To: CDC ______ |  |  |  |
| Other: |  |  |  |
| Other: |  |  |  |